*GeneLab Standard Operating Procedure: Tissue homogenization*

*using Bullet Blender Gold Bead Beater*

*Sep 2021*

*Version 1.1*

# Document Revisions

|  |  |  |  |
| --- | --- | --- | --- |
| Document Number | Revision Number | Date | Description of Changes |
| GL-SOP-2.1 | 1.0 | May 2020 | Original document |
| GL-SOP-2.1 | 1.1 | Sep 2021 | Expanded lysis buffer volume range per sample  Added names to referenced SOP’s |
|  |  |  |  |
|  |  |  |  |

# Scope and Purpose

The procedure below describes the steps required to lyse and homogenize biological material using a Bullet Blender 24 Gold bead beater. This procedure was validated for downstream RNA/DNA extraction using the Qiagen AllPrep kits (SOP #3.1), it is also possible to use the procedure with downstream RNA extraction using Trizol (SOP #3.2).

# Equipment and Consumables

1. Eppendorf Centrifuge 5424/5424 R
2. Styrofoam box and dissection tools (forceps, scalpel, scissors)
3. Bullet Blender 24 Gold (NextAdvanced Cat#BB24-AU)
4. Sterile Microcentrifuge Tube 1.5 mL RINO® (NextAdvanced Cat# TUBE1R5-S or Navy, Red, Green kits, refer to Table 1 to select appropriate tube/bead combination)

# Procedure

1. Obtain dry ice, it will be used to maintain a cold environment inside the bead beater.
2. Select the homogenization bead type, homogenization duration and homogenization speed setting based on the material processed, using table 1 below:
   1. If using beads that are not part of the kit, use RINO tubes (Next Advance #TUBE1R5-S).
   2. Kits (Navy, Red, Green) have tubes pre-filled with beads.

Table 1: Validated homogenization settings (preferred condition in bold)

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Tissue/Bead Type | NAVY kit (Next  Advanced  #NAVYR5)/Speed setting | Zirconium Oxide  2.0mm (Next  Advance  #ZROB20-  RNA)/Speed  setting | Red kit (Next  Advanced  #REDR5-  RNA)/Speed  setting | SSB (Next  Advanced  #SSB14B-RNA)/Speed setting |
| Liver | **3min/12** | 3min/12 |  |  |
| Thymus | **3min/12** |  |  |  |
| Feces | **3min/12** |  |  |  |
| Colon | **3min/12** | 5min/12 |  |  |
| Lung |  | **3min/12** |  |  |
| Skin | TBD | TBD |  |  |
| Spleen |  | **5min/12** |  |  |
| Muscle |  |  |  | **5min/12** |
| Kidney | **3min/12** |  |  |  |
| Brain | 3min/10 |  | **3min/10** |  |
| Heart |  |  |  | **5min/12** |
| Lymph Nodes |  |  | **3min/10** |  |
| White Adipose |  |  | **3min/10** |  |
| Brown Adipose |  |  | **3min/10** |  |
| Reproductive Tract | **3min/12** |  |  |  |
| Adrenal Gland |  | **3min/12** |  |  |

1. Select a lysis buffer depending on the downstream extraction protocol used:
   1. If following SOP #3.1 - Qiagen Allprep DNA\_RNA Mini with Qiagen RNase free DNase set, on the day of the extraction prepare 600-1000uL of QIAGEN buffer RLT + 1% b-ME per sample.
   2. If following SOP #3.2 - TRIzol RNA Extraction with QIAGEN RNase-Free DNase Set and QIAGEN Allprep, on the day of the extraction prepare 600-1000uL of TRIzol solution per sample.
   3. If following SOP #3.3 - Feces DNA Extraction using Maxwell RSC instrument with Purefood GMO kit, on the day of the extraction prepare 1000uL CTAB buffer per sample.
2. Pre-chill a centrifuge that can fit RINO tubes to 4°C (Suggested Eppendorf Centrifuge 5424/5424 R, this equipment will be used in the consecutive extraction SOP’s).
3. Fill the Bullet Blender Gold (BBG) side compartment with dry ice pellets or crushed dry ice.



Figure 1: Ice compartment on left side.

1. Pre-chill the bullet blender by closing the tube compartment and turning the BBG on for 5min at speed setting 12.
2. When the BBG is cold, a blue “4°C” indicator light will show on the front screen.
3. Follow SOP#1.2 for tissue cutting procedures. When completed, you should have lysis tubes with a tissue ready.
4. Close the RINO tubes **tightly** and place them inside the pre-chilled BBG. Make sure to equally distribute the weight of the RINO tubes throughout the rotor.



Figure 2: Example of a well-balanced rotor with six sample tubes (in blue).

1. Homogenize the tissues following the speed and time setting indicated in Table 1.
2. Move RINO tubes with the homogenate into the centrifuge and spin for 5min at full speed in 4°C. (Full speed using Eppendorf 5424 is 21130 x g/15000 rpm.)
3. Carefully transfer lysate in to clean, labeled 1.5ml Eppendorf tube and proceed with SOP #3.1/3.2/3.3 for RNA/DNA isolation.
4. Take out the dry ice pellets and return to storage.
5. Wipe the inside of the dry ice pocket and the tube rotor.